- Safe Handling of Cryogenic Liquids

University of California Santa Cruz

UCO/Lick Observatory – Mt. Hamilton

31 July 2026

Cryogenic Safety Training.doc

Cryogenic Safety Training.pdf

- Safe Handling of Cryogenic Liquids

Table of Contents

Introduction 	3

Hazards of Cryogenic Liquids 	3

Contact with and Destruction of living Tissue 	3

Pressure Build-Up/Explosion 	3

Flesh Tearing and Materials Embrittlement 	4

Boiling and Splashing 	4

Oxygen Enrichment 	4

Accelerating Combustion Effects of Liquid Oxygen 	4

Oxygen Deficiency/Asphyxiation 	4

Personal Protective Equipment - PPE 	5

Eye Protection 	5

Hand Protection 	5

Protective Clothing 	5

Safe Handling Procedures	6

Preparation 	6

Transfer and Use 	6

Storage 	7

First Aid 	7

Introduction

This document describes the principal hazards and appropriate safety procedures associated with two cryogenic liquids that are used at Lick Observatory: Liquid N2 and Helium. The safe handling of cryogenic liquids involves understanding the unique properties of these materials and ensuring that appropriate safety precautions are taken at all times.

Cryogenic liquids are typically odorless and colorless when vaporized to the gaseous state.  Most of them have no color as a liquid although liquid oxygen is light blue.  However, the extremely cold liquid and vapor have a built in warning property as the cold boil off gases condense moisture in the air, creating a highly visible fog.

Hazards of Cryogenic Liquids

Contact With and Destruction of Living Tissues

Since all cryogenic liquids and their cold boil-off vapors are extremely cold they can 
rapidly freeze human tissue and cause frostbite.  Even a brief contact with a cryogenic liquid is capable of causing tissue damage similar to that of thermal burns. Prolonged contact may result in blood clots.

Pressure Build-Up/Explosions

Cryogenic liquids exhibit large volume exchange ratios which can cause rapid pressure 
changes.  One volume of liquid nitrogen, for example, will vaporize to approximately 
700 volumes of nitrogen gas when warmed to room temperature at one atmosphere. In addition, all cryogens can condense sufficient moisture from the air to block the 
opening in storage vessels.  This condition can result in an explosion caused by the 
build up of trapped vapor in the container.  As a result, it is critical to contain these 
liquids in insulated dewars with pressure relief valves (see "Safe Handling Procedures" 
below).

In case of the uncontrolled release of cryogens:

1. Leave the area immediately

2. Never enter a vapor cloud.

- 3. If you find yourself in the path of a vapor cloud, hold your breath until you are out of the cloud.

Flesh Tearing and Materials Embrittlement

The touching of uninsulated containers or other materials that have been cooled by 
cryogenic liquids can cause serious skin injuries.  The extremely cold surface of the 
cooled material will cause the flesh to stick fast and tear when one attempts to 
withdraw from it.  Even non-metallic materials are dangerous to touch at low 
temperatures.  In addition to the hazards of frostbite or flesh sticking to cold materials, objects that are soft and pliable at room temperature, are easily broken because they become hard and brittle at extremely low temperatures and will easily fracture.

Boiling and Splashing

Cryogenic liquids can boil or splash when added to a warm container or when inserting 
warm objects in the liquid.  Always perform these operations slowly to minimize boiling and splashing.

Oxygen Enrichment

Cryogenic liquids with a boiling point below that of liquid oxygen (i.e. nitrogen and 
helium) have the ability to condense oxygen out of the air if exposed to the atmosphere.  Violent reactions (e.g. rapid combustion or explosions) may occur if the system or process is not compatible with liquid oxygen (see below).

Accelerating Combustion Effect of Liquid Oxygen

Oxygen is not flammable, but it vigorously accelerates and supports combustion. 
Substances that burn in air will burn more vigorously in oxygen.  Do not permit liquid 
oxygen or oxygen rich atmosphere to come into contact with organic materials or 
flammable substances of any kind.  Some organic materials which can react violently 
with liquid oxygen include oil, grease, asphalt, kerosene, cloth, tar and dirt.

Oxygen Deficiency/Asphyxiation

All cryogenic liquids have a significant potential for creating an oxygen deficient 
environment.  Because of their large liquid to gas expansion ratios, even small 
quantities of liquefied gas can expand to displace large amounts of oxygen, thereby 
rendering the atmosphere in a confined area lethal.  Without adequate oxygen, 
one can lose consciousness in a few seconds and die of asphyxiation in a few 
minutes.

Personal Protective Equipment - PPE

Eye Protection

Safety glasses must be worn at all times while handling cryogenic liquids, because 
the liquid is almost always boiling and can splash into the eyes.  Safety goggles 
provide the best protection for the eyes. When filling dewars or transferring 
cryogenic liquids from one container to another, face shields must also be worn.

Hand Protection

When cryogenic liquids are spilled on the skin, a thin gaseous layer apparently forms 
next the skin. This thin layer protects the skin from freezing, provided the contact 
with the cryogen is brief and in small quantities.

The most likely cause of frostbite to the hands and body is by contact with cold 
metal surfaces. Because there is no protective layer of gas formed, frostbite will 
occur almost instantaneously when touching surfaces that have been cooled by 
cryogenic liquids.   Hand protection is primarily required to guard against the hazard 
of touching cold surfaces.  Loose, non-asbestos insulating gloves that can be easily removed may be worn. Never use rags to handle cold surfaces. Use tongs to add or remove materials from cryogenic liquids.

Protective Clothing

Closed toed shoes are required when handling cryogenic liquids.  Long cuffless 
trousers which cover the neck of the shoes should also be worn.  Long sleeve 
shirts are also recommended for arm protection.  An apron made of leather or 
other material recommended for use with cryogens is indicated when large 
quantities of cryogen are handled. A cryo-apron is available in the downstairs cryo-transfer area.

Safe Handling Procedures

Preparation

1. Always be familiar with the hazards of the liquid in use.

2. Work in an open, well-ventilated location.

3. Ensure that safety glasses and if necessary face shields are worn.

4. Have appropriate gloves and apron on hand.

5. Examine containers and pressure relief valves for signs of defect. Never use a

container which has defects.

6. Any exposed glass areas of dewars should be taped to prevent the generation
    of broken glass projectiles in the event the container implodes.

7. Ensure that all equipment and containers are free of oil, grease, dirt or other materials which may create a hazardous condition upon contact with the cryogenic   liquid.  Employees should wash their hands and arms with soap and water, rinsing and drying thoroughly prior to handling the liquid after being exposed to lubricants. Clothing should be relatively clean.

8. Select working materials carefully.  Cryogenic temperatures may alter the physical

characteristics of many materials.

Transfer and Use

1. Use only fitted transfer tubes designed for use with the dewar container. 
    Damaged transfer tubes should be replaced.  Do not handle transfer tubes 
    with bare hands as the fitting is not insulated.

2. When transferring into a secondary container, do not fill the secondary container 
    to more than 80% of capacity (60% if the temperature is likely to be above 30 C).

3. Do not lower experiments into storage dewars unless provisions have been made 
    to vent the dewar and prevent freezing in the narrow neck.

4. Immediately re-cap any container to prevent atmospheric moisture from entering 
    and forming an ice plug.

5. Provide proper venting for the dewars used in experiments.

6. Use care in transporting fragile cryogenic containers.  Use a hand truck for 
    transport.  Always transport cryogenic liquids in service elevators when available.

- Storage

1. Store in a well ventilated area to prevent buildup of flammable gases or air 
    displacement.

2. Use only approved storage vessels having pressure relief valves.

3. Avoid contact of moisture with storage containers to prevent ice plugging 
    of relief devices.

4. Periodically check container necks for ice plugging; core out plugs if present.

5. Keep all sources of ignition away from cryogenic liquids.

- First Aid

If medical attention is needed following a skin or eye exposure to a cryogenic 
liquid, immediately call 5911.  All personnel using cryogenic liquids must also be familiar with first aid procedures for treating frostbite:

1. Warm the affected area of the body rapidly by immersion in water not to exceed 
105 F, with body heat, or by exposure to warm air. Maintain the affected area of the victim at normal body temperature until medical help arrives.

2. Calm the victim and avoid aggravating injury.  People with frostbitten feet 
    should not walk on them.  Do not rub or massage the affected parts of the body.

3. Prevent infection--use a mild soap to clean the affected area. 
    Dressings need not be applied if the skin is intact.

4. If affected, flush eyes with warm water for 15 minutes.